## MainTable
Results Data in [Sample Results] for Test Name | LC-MS
Analyte | Acrylamide
Report Type (Spike, Blank, Ref…) | Low Spike
Units | ng/L
Low Spike Concentration | 90
MDL (From Mean of Previous Low Spikes)
Number of Significant figures | 2
One or two tail Grubbs’ Test (α: 0.05) | 1
Total Number of Results | 9 (1 Excluded)
Mean of Samples (Ymv) | 83.168
ST DEV σ of Results | 6.377
RSD% (CV) | 7.668%
bias as %age (100 - Recovery) | 7.592%
MDL Factor (n-1 Degrees Freedom) | 2.998
Units - 1 Part in This Value | 1000000000000
Predicted Horwitz "Expected Value" as RSD% | 22
Horwitz Ratio (Ideally 0.3≤ HorRat ≤1) | 0.34853
Ymax (Maximum Value) | 89.84
Ymin (Minimum Value) | 72.11
Result to 2 Significant figures | 83
MDL = (MDL Factor x SD) | 19
PQL (Practical Quantitation limit) | 57
% Recovery | 92.41%
Number of Results included in Calculations | 8
Ymax - Ymv | 6.6725
Ymv - Ymin | 11.0575
Sig Value: 1-sided | 0.0125
G (min) | 1.734
G (max) | 1.0463
G Critical | 2.032
Check if This Result is an Outlier | 0
Number of Excluded Results | 1
Number of Decimal places for DISPLAYED Data | 3
Raw Mean of Samples (Ymv) | 83.1675
Uncertainty% at 95% | 21.58%
Raw ST DEV | 6.3775
Raw Relative SD | 0.076677
Raw uncertainty at 95Conf - u | 6.7638
Raw MDL = (MDL Factor x SD) | 19.1182
Raw PLQ | 57.3547
Raw Bias | 13.5277
Raw Spike Value | 90
Raw bias (Mean - Spike) | -6.832
MDL Corrected | 20.6889
MDL rounded | 19
PQL rounded | 57
Spike Level to correct Sig. Figs | 90
Raw bias corrected for recovery | 14.639
MDL Corrected rounded | 21
Corrected for Recovery Check [Report] Sheet >> | 90 ± 21 ng/L

## SampleResults
S1 | 86.38
S2 | 87.79
S3 | 77.94
S4 | 72.11
S5 Duplicate S3 | 78.56
S6 | 89.84
S7 | 89.3
S8 | 83.42
S9 (Outlier?) | !59.97

## Report
Summary Report - Low Spike
Report Date | 2022-04-25
Laboratory Name | Tim’s Testing & Consulting Laboratory Pty Ltd
Address | 15a Bridge Street
Some-Town, STATE, X100
Contact | Tim Strutt
e.g. tcstrutt+horwitzt@gmail.com
Authorised By | Saga Brondam
Analyte | Acrylamide
Number of Samples in This Report | 8
Report Type | Low Spike
Low Spike Level | 90 ng/L
% Recovery | 92.41%
RSD% (CV) | 0.076677
Bias as a %age | 0.075917
Is Bias Satisfactory? Y/N | Yes
Horwitz Ratio (Ideally 0.3≤ HorRat ≤1) | 0.34853
Is Horwitz Ratio Satisfactory? (Y/N): | Yes
Result to 2 Significant figures | 83 ng/L
Method Detection Limit (MDL) | 19 ng/L
Practical Quantitation Limit (PQL) | 57 ng/L
Result Corrected for Recovery Bias | 90 ± 21 ng/L

## Distributions
Student’s t-test |  | Grubbs’ test 𝜶 0.05
Limit | 99% | 1-sided | 2-sided
1 | 31.82
2 | 6.965
3 | 4.541 | 1.153 | 1.155
4 | 3.747 | 1.463 | 1.481
5 | 3.365 | 1.672 | 1.715
6 | 3.143 | 1.822 | 1.887
7 | 2.998 | 1.938 | 2.02
8 | 2.896 | 2.032 | 2.126
9 | 2.821 | 2.11 | 2.215
10 | 2.764 | 2.176 | 2.29
11 | 2.718 | 2.234 | 2.355
12 | 2.681 | 2.285 | 2.412
13 | 2.65 | 2.331 | 2.462
14 | 2.624 | 2.371 | 2.507
15 | 2.602 | 2.409 | 2.549
16 | 2.583 | 2.443 | 2.585
17 | 2.567 | 2.475 | 2.62
18 | 2.552 | 2.504 | 2.651
19 | 2.539 | 2.532 | 2.681
20 | 2.528 | 2.557 | 2.709
21 | 2.518 | 2.58 | 2.733
22 | 2.508 | 2.603 | 2.758
23 | 2.5 | 2.624 | 2.781
24 | 2.492 | 2.644 | 2.802
25 | 2.485 | 2.663 | 2.822
26 | 2.479 | 2.681 | 2.841
27 | 2.473 | 2.698 | 2.859
28 | 2.467 | 2.714 | 2.876
29 | 2.462 | 2.73 | 2.893
30 | 2.457 | 2.745 | 2.908
40 | 2.423 | 2.866 | 3.036
50 | 2.403 | 2.956 | 3.128
60 | 2.39 | 3.025 | 3.199
80 | 2.374 | 3.13 | 3.3105
100 | 2.364 | 3.207 | 3.383
150 | 2.352 | 3.35 | 3.515
200 | 2.345 | 3.5 | 3.7
1000 | 2.33 | 4 | 4.5
Student’s t-distributions have ν degrees of freedom for regions.
See www.itl.nist.gov/div898/handbook/, 2012.
Horwitz Ratio - Ref: RSC, M Thompson AMC Technical Brief No.17 July 2004
https://www.rsc.org/images/horwitz-function-technical-brief-17_tcm18-214859.pdf
Grubbs’ statistic values are at the 0.05 (95%) level. Here for <9 results 1-tail is recommended. If there are more results, a 2-tail may be preferable (one outlier in each tail). To remove a result, add text (e.g. 123.4?). Removing 2 results from data sets of <7 is not recommended. The method is suspect if outliers are in the same tail.
Ref: RSC - Analytical Methods Committee, AMCTB No. 69
https://pubs.rsc.org/en/content/articlehtml/2015/ay/c5ay90053k and
Frank E. Grubbs and Glenn Beck, "Extension of Sample Sizes and Percentage Points for Significance Tests of Outlying Observations", Technometrics, 14(4), 847-854 (1972).

## Units
Unit | NumericValue
ppm | 1000000
ppb | 1000000000
ppt | 1000000000000
mg/kg | 1000000
µg/kg | 1000000000
ng/kg | 1000000000000
mg/L | 1000000
µg/L | 1000000000
ng/L | 1000000000000
pg/L | 1000000000000000
g/L | 100
percent | 100
drops/floz | 568
If a Unit is added, changed, or deleted - Check that changes to the Unit are in the [MainTable] Select Units from List Pop-Up Menu. If there is a mismatch Error Warnings will be displayed…

## Notes
This spreadsheet may be useful to assist Laboratory Analysts and Assessors ensure compliance with ISO/ IEC 17025:2017. The revised standard requires Measurement Uncertainty (MU) of chemical test results to be available. It is designed to help produce necessary information by using sample blanks, spikes, and comparisons with reference materials. Sample "Low Spikes" could be used to determine Recoveries and Method Detection Limits. Spikes establish MUs at different concentrations with different matrices. It may be helpful when developing new methods or for "novel" analytes and matrices. An authorised report for a single analyte can be printed.
Created with Apple "Numbers" Version 12.0, and Microsoft "Excel" 2013, spreadsheet programs - Using only "Cell Functions". No "Code", Macros, or Hyperlinks are used - It should be “Safe” to distribute…
The spreadsheet uses some conventions: In [MainTable] cells where the user is expected to input data have blue text on a white background, bold blue indicates that a selection is made from a list. Default look-ups are grey. Warnings are highlighted with an [orange background]; and critical/failure/out-of-range warnings are highlighted with [red backgrounds]. Important data is purple.
Getting Started: The sheets are [MainTable], [SampleResults], [Report], [Distributions] and [Units], and [Notes]. The last three are "look-ups" or information. Data should be added or changed as outlined below…
[MainTable]: Used to set the parameters of the Test:- Fill in the Instrument/Test name that was used to create the results. Fill in the Analyte Name for your set of results Select Report Type from the drop-down - This will produce different a different [Report] for each type. Select Units from the drop-down - This will change the Horwitz Ratio (larger concentrations should give higher precision). Enter 'Spike Value' - Values between 0.00001 and 10,000,000. MDL (From Mean of Previous Low Spikes) - If a value is entered, it will be used instead of the Results calculation. Select Number of Significant figures - Usually 2 - 1 may be useful if near Quantitation limits, 3 for special precision. Select One or two tail Grubbs’ Test (α: 0.05) - Normally 1; 2 is not recommended for <9 approved Results. 1 calculates the right and left tails separately, the furthest from the mean is used for the "one outlier" hypothesis. Now go to the [SampleResults] Sheet… [SampleResults]: Used to add SampleID and Results. There are 2 columns:- Used to Add/Edit SampleID (or Paste from another spreadsheet). Used to Add/Edit Results (or Paste from another spreadsheet). Possible Outliers can be indicated by adding text e.g. Changing 77 to ?77 or 77! This will cause the possible Outlier to be ignored by the Horwitz and Grubbs calculations. The sheet can have many rows, but normally 3-11 are used (5, or more recommended). [Report]: Used to display or print a formal report:- The Report Date, Laboratory Name, Address, Contact, and Authorised data can be edited. Check that the Bias and Horwitz Ratio are satisfactory - An Outlier can increase these values… [Distributions]: Used to calculate Method Detection Limit (t-test) and Grubbs’ critical value:- This is not editable. It is recommended that the user consults the NIST and RSC papers, referenced in the text at the bottom of the sheet, for further background information. [Units]: Used to calculate scale of Horwitz values, and in [Report] and [Calculations - Main Table]:- It is normally Locked, but any changes must be reflected in [MU Calculations] "Select Units" drop-down.
For further information, please consult these References:- Planning and Reporting Method Validation Studies; Supplement to EURACHEM Guide on the Fitness for Purpose of Analytical Methods, First edition 2019 Blanks in Method Validation, Supplement to EURACHEM Guide The Fitness for Purpose of Analytical Methods, First Edition 2019
The password to unlock individual MS Excel sheets and cells is "timstrutt" (no "quotes"). Apple Sheets use Locked "Shapes".
The [MainTable] uses "hidden" rows with a red font - Changing or deleting them will cause breakages...
Warning: Some of this content is the author’s opinion. The calculations are based upon standard literature methods, but some approximations have been used. See the references described above for background information.
This spread-sheet is "Public Domain". Anyone is free to copy, modify, publish, use, or distribute it for any purpose, commercial or non-commercial, and by any means. The author is retired; carries out "Pro Bono" work only; and does not accept any responsibility for any loss caused by the use of, or failure to use, this spreadsheet. Please do not remove my name, copyright & version information.
Version: 20220915 - ©Tim Strutt.