## MainTable
Test Name (TopRow, 1st Column in SampleResults) |  | ICP-MS | ICP-MS | ICP-MS | ICP-MS | ICP-MS | ICP-MS | ICP-MS | ICP-MS | ICP-MS | ICP-MS | ICP-MS | ICP-MS | ICP-MS | ICP-MS | ICP-MS | ICP-MS | ICP-MS | ICP-MS | ICP-MS | ICP-MS | ICP-MS | ICP-MS | ICP-MS | ICP-MS
Analyte (Top Row in SampleResults Sheet) |  | Al | As | B | Ba | Be | Cd | Co | Cr | Cu | Fe | Hg | Mn | Mo | Ni | Pb | Sb | Se | Sn | Sr | Ti | V | Zn | Rh | Na
Select Report Type (Spike, Ref…) : | Low Spike | Low Spike | Low Spike | Low Spike | Low Spike | Low Spike | Low Spike | Low Spike | Low Spike | Low Spike | Low Spike | Low Spike | Low Spike | Low Spike | Low Spike | Low Spike | Low Spike | Low Spike | Low Spike | Low Spike | Low Spike | Low Spike | Low Spike | Low Spike | Low Spike
Select Units (ppm, µg/kg, g/L…) : | µg/kg | µg/kg | µg/kg | µg/kg | µg/kg | µg/kg | µg/kg | µg/kg | µg/kg | µg/kg | µg/kg | µg/kg | µg/kg | µg/kg | µg/kg | µg/kg | µg/kg | µg/kg | µg/kg | µg/kg | µg/kg | µg/kg | µg/kg | µg/kg | µg/kg
Select # of Significant Figures : | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2
Enter Expected Analyte Concentration For Analytes : |  | 90 | 2.2 | 99 | 90 | 92 | 9 | 95 | 90 | 90 | 90 | 1 | 90 | 102 | 99 | 90 | 93 | 108 | 90 | 0.00101 | 106 | 1.155 | 0.1155 | 0.01155 | 1151
MDL - Optional (From Mean of Previous Low Spikes) : | bar: Fe
Total Number of Results for Analyte |  | 9 | 9 | 9 | 9 | 9 | 9 | 9 | 9 | 9 | 9 | 9 | 9 | 9 | 9 | 9 | 9 | 9 | 9 | 9 | 9 | 9 | 9 | 9 | 9
RSD% (CV) |  | 12.47% | 8.758% | 4.172% | 16.28% | 143.7% | 6.907% | 7.086% | 14.36% | 7.43% | 6.716% | 7.409% | 9.747% | 6.897% | 7.246% | 11.2% | 10.1% | 7.199% | 7.822% | 10.86% | 9.479% | 10.92% | 10.92% | 10.95% | 4.801%
bias as %age (100 - Recovery) |  | 10.58% | 26.97% | 5.688% | -12.36% | -73.47% | -3.259% | -2.105% | 11.22% | -3.457% | -10.11% | 16.37% | -8.889% | 4.357% | 16.24% | -9.753% | -9.08% | 23.11% | -11.85% | -22.17% | 23.08% | -6.667% | -6.667% | -6.59% | 12.72%
Horwitz Ratio (Ideally 0.3≤HorRat≤1) Check: Very High |  | 0.56662 | 0.39809 | 0.18965 | 0.74006 | 6.5324 | 0.31396 | 0.32211 | 0.65264 | 0.33773 | 0.30529 | 0.33679 | 0.44307 | 0.31352 | 0.32938 | 0.5092 | 0.4592 | 0.32725 | 0.35554 | 0.49353 | 0.43085 | 0.49615 | 0.49615 | 0.49791 | 0.30863
Grubbs' Outlier Check : 4 Found? |  | 58.97 | 0 | 0 | 140.1 | 771 | 0 | 0 | 53.77 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Result to 2 Significant figures |  | 80 | 1.6 | 93 | 100 | 160 | 9.3 | 97 | 80 | 93 | 99 | 0.84 | 98 | 98 | 83 | 99 | 100 | 83 | 100 | 0.0012 | 82 | 1.2 | 0.12 | 0.012 | 1000
MDL = (MDL Factor x SD) |  | 32 | 0.6 | 12 | 42 | 383 | 1.8 | 19 | 37 | 19 | 1 | 0.2 | 25 | 20 | 21 | 29 | 27 | 20 | 20 | 0.0003 | 30 | 0.4 | 0.04 | 0.004 | 200
PQL (Practical Quantitation limit) |  | 97 | 1.7 | 36 | 130 | 1100 | 5.4 | 58 | 110 | 58 | 2.7 | 0.64 | 76 | 61 | 62 | 88 | 82 | 68 | 61 | 0.00095 | 87 | 1.1 | 0.11 | 0.011 | 480
% Recovery |  | 89.42% | 73.03% | 94.31% | 112.4% | 173.5% | 103.3% | 102.1% | 88.78% | 103.5% | 110.1% | 83.63% | 108.9% | 95.64% | 83.76% | 109.8% | 109.1% | 76.89% | 111.9% | 122.2% | 76.92% | 106.7% | 106.7% | 106.6% | 87.28%
Is Result is an Outlier (>0)? |  | 58.97 | 0 | 0 | 140.1 | 771 | 0 | 0 | 53.77 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Ymax (Maximum Value) |  | 89.84 | 1.82 | 99.21 | 140.1 | 771 | 10.4 | 106 | 89.84 | 103 | 111.3 | 0.898 | 117 | 106 | 89.84 | 121 | 122 | 89.84 | 117 | 0.00146 | 89.84 | 1.46 | 0.146 | 0.0146 | 1100
Ymin (Minimum Value) Outlier? |  | 58.97 | 1.36 | 85.37 | 81.31 | 72.11 | 8.07 | 83 | 53.77 | 84 | 90.1 | 0.721 | 90 | 89 | 72.11 | 90 | 84 | 72.11 | 88 | 0.00103 | 68.43 | 1.03 | 0.103 | 0.0103 | 951
Result Corrected for Recovery |  | 90 | 2.2 | 99 | 90 | 92 | 9.0 | 95 | 90 | 90 | 90 | 1.0 | 90 | 100 | 99 | 90 | 93 | 110 | 90 | 0.0010 | 110 | 1.2 | 0.12 | 0.012 | 1200
Number of Results included in Calculations |  | 9 | 9 | 9 | 9 | 9 | 9 | 9 | 9 | 9 | 9 | 9 | 9 | 9 | 9 | 9 | 9 | 9 | 9 | 9 | 9 | 9 | 9 | 9 | 9
Raw Mean of Samples (Ymv) |  | 80.4789 | 1.6067 | 93.3689 | 101.1211 | 159.5933 | 9.2933 | 97 | 79.9011 | 93.1111 | 99.1 | 0.83633 | 98 | 97.5556 | 82.9267 | 98.7778 | 101.4444 | 83.0378 | 100.6667 | 0.0012339 | 81.53 | 1.232 | 0.1232 | 0.012311 | 1004.5556
Raw ST DEV σ of Results |  | 10.0322 | 0.14071 | 3.8956 | 16.4639 | 229.3551 | 0.64189 | 6.8739 | 11.4722 | 6.9182 | 6.6558 | 0.061968 | 9.5525 | 6.7289 | 6.0092 | 11.0655 | 10.2483 | 5.9783 | 7.874 | 0.00013397 | 7.7279 | 0.13448 | 0.013448 | 0.0013486 | 48.2263
MDL Factor (n-1 Degrees Freedom) |  | 2.896 | 2.896 | 2.896 | 2.896 | 2.896 | 2.896 | 2.896 | 2.896 | 2.896 | 2.896 | 2.896 | 2.896 | 2.896 | 2.896 | 2.896 | 2.896 | 2.896 | 2.896 | 2.896 | 2.896 | 2.896 | 2.896 | 2.896 | 2.896
Units - 1 Part in This Value |  | 1000000000 | 1000000000 | 1000000000 | 1000000000 | 1000000000 | 1000000000 | 1000000000 | 1000000000 | 1000000000 | 1000000000 | 1000000000 | 1000000000 | 1000000000 | 1000000000 | 1000000000 | 1000000000 | 1000000000 | 1000000000 | 1000000000 | 1000000000 | 1000000000 | 1000000000 | 1000000000 | 1000000000
Predicted Horwitz "Expected Value" as RSD% |  | 22 | 22 | 22 | 22 | 22 | 22 | 22 | 22 | 22 | 22 | 22 | 22 | 22 | 22 | 22 | 22 | 22 | 22 | 22 | 22 | 22 | 22 | 22 | 15.555
Ymax - Ymv |  | 9.3611 | 0.21333 | 5.8411 | 38.9789 | 611.4067 | 1.1067 | 9 | 9.9389 | 9.8889 | 12.2 | 0.061667 | 19 | 8.4444 | 6.9133 | 22.2222 | 20.5556 | 6.8022 | 16.3333 | 0.00022611 | 8.31 | 0.228 | 0.0228 | 0.0022889 | 95.4444
Ymv - Ymin |  | 21.5089 | 0.24667 | 7.9989 | 19.8111 | 87.4833 | 1.2233 | 14 | 26.1311 | 9.1111 | 9 | 0.11533 | 8 | 8.5556 | 10.8167 | 8.7778 | 17.4444 | 10.9278 | 12.6667 | 0.00020389 | 13.1 | 0.202 | 0.0202 | 0.0020111 | 53.5556
Sig Value: 1-sided |  | 0.011111 | 0.011111 | 0.011111 | 0.011111 | 0.011111 | 0.011111 | 0.011111 | 0.011111 | 0.011111 | 0.011111 | 0.011111 | 0.011111 | 0.011111 | 0.011111 | 0.011111 | 0.011111 | 0.011111 | 0.011111 | 0.011111 | 0.011111 | 0.011111 | 0.011111 | 0.011111 | 0.011111
G (min) |  | 2.144 | 1.753 | 2.0533 | 1.2033 | 0.38143 | 1.9058 | 2.0367 | 2.2778 | 1.317 | 1.3522 | 1.8612 | 0.83748 | 1.2715 | 1.8 | 0.79326 | 1.7022 | 1.8279 | 1.6087 | 1.5219 | 1.6952 | 1.5021 | 1.5021 | 1.4913 | 1.1105
G (max) |  | 0.9331 | 1.5161 | 1.4994 | 2.3675 | 2.6658 | 1.7241 | 1.3093 | 0.86634 | 1.4294 | 1.833 | 0.99514 | 1.989 | 1.255 | 1.1505 | 2.0083 | 2.0058 | 1.1378 | 2.0743 | 1.6877 | 1.0753 | 1.6955 | 1.6955 | 1.6973 | 1.9791
G Critical |  | 2.11 | 2.11 | 2.11 | 2.11 | 2.11 | 2.11 | 2.11 | 2.11 | 2.11 | 2.11 | 2.11 | 2.11 | 2.11 | 2.11 | 2.11 | 2.11 | 2.11 | 2.11 | 2.11 | 2.11 | 2.11 | 2.11 | 2.11 | 2.11
Number of Excluded Results |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Raw uncertainty at 95Conf - u |  | 10.5749 | 0.14832 | 4.1063 | 17.3545 | 241.7615 | 0.67661 | 7.2457 | 12.0928 | 7.2924 | 7.0159 | 0.06532 | 10.0692 | 7.0929 | 6.3342 | 11.664 | 10.8027 | 6.3016 | 8.2999 | 0.00014122 | 8.1459 | 0.14175 | 0.014175 | 0.0014215 | 50.835
Raw MDL = (MDL Factor x SD) |  | 29.0534 | 0.4075 | 11.2817 | 47.6796 | 664.2124 | 1.8589 | 19.9067 | 33.2235 | 20.035 | 1 | 0.17946 | 27.664 | 19.4868 | 17.4026 | 32.0456 | 29.6791 | 17.313 | 22.8031 | 0.00038798 | 22.3801 | 0.38944 | 0.038944 | 0.0039054 | 139.6634
Bias |  | 21.1498 | 0.29665 | 8.2127 | 34.709 | 483.523 | 1.3532 | 14.4914 | 24.1856 | 14.5848 | 14.0317 | 0.13064 | 20.1384 | 14.1857 | 12.6685 | 23.328 | 21.6053 | 12.6033 | 16.5999 | 0.00028244 | 16.2919 | 0.2835 | 0.02835 | 0.002843 | 101.67
Raw bias (Mean - Spike) |  | -9.5211 | -0.59333 | -5.6311 | 11.1211 | 67.5933 | 0.29333 | 2 | -10.0989 | 3.1111 | 9.1 | -0.16367 | 8 | -4.4444 | -16.0733 | 8.7778 | 8.4444 | -24.9622 | 10.6667 | 0.00022389 | -24.47 | 0.077 | 0.0077 | 0.00076111 | -146.4444
MDL Corrected |  | 32.4905 | 0.55799 | 11.9621 | 42.4358 | 382.8953 | 1.8002 | 19.4963 | 37.4228 | 19.3656 | 0.90817 | 0.21458 | 25.4057 | 20.3746 | 20.7756 | 29.1979 | 27.2085 | 22.5176 | 20.3869 | 0.00031758 | 29.0971 | 0.3651 | 0.03651 | 0.003664 | 160.0236
PLQ Corrected |  | 97.4716 | 1.674 | 35.8863 | 127.3075 | 1148.6859 | 5.4007 | 58.4888 | 112.2683 | 58.0968 | 2.7245 | 0.64373 | 76.2171 | 61.1238 | 62.3269 | 87.5936 | 81.6256 | 67.5527 | 61.1607 | 0.00095275 | 87.2913 | 1.0953 | 0.10953 | 0.010992 | 480.0708
PQL rounded |  | 97 | 1.7 | 36 | 130 | 1100 | 5.4 | 58 | 110 | 58 | 2.7 | 0.64 | 76 | 61 | 62 | 88 | 82 | 68 | 61 | 0.00095 | 87 | 1.1 | 0.11 | 0.011 | 480
Spike Level to correct Sig. Figs |  | 90 | 2.2 | 99 | 90 | 92 | 9.0 | 95 | 90 | 90 | 90 | 1.0 | 90 | 100 | 99 | 90 | 93 | 110 | 90 | 0.0010 | 110 | 1.2 | 0.12 | 0.012 | 1200
Now Calculate correct rounding for MDL:-
Get int/log factor for Corrected Spike |  | 1 | 0 | 1 | 1 | 1 | 0 | 1 | 1 | 1 | 1 | 0 | 1 | 2 | 1 | 1 | 1 | 2 | 1 | -3 | 2 | 0 | -1 | -2 | 3
Get int/log factor for Corrected MDL |  | 1 | -1 | 1 | 1 | 2 | 0 | 1 | 1 | 1 | -1 | -1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | -4 | 1 | -1 | -2 | -3 | 2
Calculate rounding integer for MDL |  | 1 | 2 | 1 | 1 | 0 | 1 | 1 | 1 | 1 | 3 | 2 | 1 | 2 | 1 | 1 | 1 | 2 | 1 | 2 | 2 | 2 | 2 | 2 | 2
Round Corrected MDL to create correct rounding |  | 32 | 0.6 | 12 | 42 | 383 | 1.8 | 19 | 37 | 19 | 1 | 0.2 | 25 | 20 | 21 | 29 | 27 | 20 | 20 | 0.0003 | 30 | 0.4 | 0.04 | 0.004 | 200
Corrected for Recovery with Uncertainty |  | 90 ± 32 µg/kg | 2.2 ± 0.6 µg/kg | 99 ± 12 µg/kg | 90 ± 42 µg/kg | 92 ± 383 µg/kg | 9.0 ± 1.8 µg/kg | 95 ± 19 µg/kg | 90 ± 37 µg/kg | 90 ± 19 µg/kg | 90 ± 1 ^µg/kg | 1.0 ± 0.2 µg/kg | 90 ± 25 µg/kg | 100 ± 20 µg/kg | 99 ± 21 µg/kg | 90 ± 29 µg/kg | 93 ± 27 µg/kg | 110 ± 20 µg/kg | 90 ± 20 µg/kg | 0.0010 ± 0.0003 µg/kg | 110 ± 30 µg/kg | 1.2 ± 0.4 µg/kg | 0.12 ± 0.04 µg/kg | 0.012 ± 0.004 µg/kg | 1200 ± 200 µg/kg

## SampleResults
ICP-MS | Al | As | B | Ba | Be | Cd | Co | Cr | Cu | Fe | Hg | Mn | Mo | Ni | Pb | Sb | Se | Sn | Sr | Ti | V | Zn | Rh | Na
S1 | 86.38 | 1.59 | 94.21 | 94.21 | 86.38 | 9.03 | 96 | 86.38 | 90 | 96.1 | 0.851 | 90 | 90 | 86.38 | 90 | 84 | 86.38 | 96 | 0.00103 | 86.38 | 1.03 | 0.103 | 0.0103 | 962
S2 | 87.79 | 1.66 | 91.3 | 88.46 | 87.79 | 9.11 | 97 | 87.79 | 84 | 90.1 | 0.882 | 91 | 104 | 87.79 | 91 | 97 | 87.79 | 97 | 0.00122 | 87.79 | 1.2 | 0.12 | 0.012 | 973
S3 | 77.94 | 1.42 | 91.88 | 98.27 | 77.94 | 9.24 | 98 | 77.94 | 101 | 95.2 | 0.779 | 92 | 93 | 77.94 | 92 | 98 | 77.94 | 88 | 0.00125 | 77.94 | 1.25 | 0.125 | 0.0125 | 981
S4 | 72.11 | 1.59 | 95.31 | 99.41 | 72.11 | 9.35 | 83 | 72.11 | 93 | 91.3 | 0.721 | 93 | 89 | 72.11 | 93 | 99 | 72.11 | 99 | 0.00133 | 72.11 | 1.33 | 0.133 | 0.0133 | 989
S5 Duplicate S3 | 78.56 | 1.66 | 94.69 | 100.24 | 78.56 | 9.06 | 100 | 78.56 | 85 | 100.2 | 0.786 | 94 | 100 | 78.56 | 94 | 100 | 78.56 | 100 | 0.00146 | 78.56 | 1.46 | 0.146 | 0.0146 | 951
S6 | 89.84 | 1.36 | 85.37 | 100.91 | 89.84 | 9.47 | 101 | 89.84 | 95 | 111.3 | 0.898 | 95 | 91 | 89.84 | 95 | 101 | 89.84 | 101 | 0.00137 | 89.84 | 1.371 | 0.1371 | 0.0137 | 1005
S7 | 89.3 | 1.69 | 99.21 | 81.31 | 89.3 | 9.91 | 102 | 89.3 | 99 | 102.4 | 0.893 | 99 | 102 | 89.3 | 99 | 102 | 89.3 | 102 | 0.00114 | 89.3 | 1.138 | 0.1138 | 0.0114 | 1024
S8 | 83.42 | 1.82 | 96.33 | 107.18 | 83.42 | 8.07 | 90 | 83.42 | 103 | 101.6 | 0.833 | 117 | 103 | 83.42 | 121 | 110 | 83.42 | 106 | 0.001135 | 83.42 | 1.135 | 0.1135 | 0.0113 | 1056
S9 | 58.97 | 1.67 | 92.02 | 140.1 | 771 | 10.4 | 106 | 53.77 | 88 | 103.7 | 0.884 | 111 | 106 | 81 | 114 | 122 | 82 | 117 | 0.00117 | 68.43 | 1.174 | 0.1174 | 0.0117 | 1100

## Distributions
Student’s t-test |  | Grubbs’ test
Limit | 99% | 𝜶 0.05 1-sided
1 | 31.82
2 | 6.965
3 | 4.541 | 1.153
4 | 3.747 | 1.463
5 | 3.365 | 1.672
6 | 3.143 | 1.822
7 | 2.998 | 1.938
8 | 2.896 | 2.032
9 | 2.821 | 2.11
10 | 2.764 | 2.176
11 | 2.718 | 2.234
12 | 2.681 | 2.285
13 | 2.65 | 2.331
14 | 2.624 | 2.371
15 | 2.602 | 2.409
16 | 2.583 | 2.443
17 | 2.567 | 2.475
18 | 2.552 | 2.504
19 | 2.539 | 2.532
20 | 2.528 | 2.557
21 | 2.518 | 2.58
22 | 2.508 | 2.603
23 | 2.5 | 2.624
24 | 2.492 | 2.644
25 | 2.485 | 2.663
26 | 2.479 | 2.681
27 | 2.473 | 2.698
28 | 2.467 | 2.714
29 | 2.462 | 2.73
30 | 2.457 | 2.745
40 | 2.423 | 2.866
50 | 2.403 | 2.956
60 | 2.39 | 3.025
80 | 2.374 | 3.13
100 | 2.364 | 3.207
150 | 2.352 | 3.35
200 | 2.345 | 3.5
1000 | 2.33 | 4
Student’s t-distributions have ν degrees of freedom for regions.
See www.itl.nist.gov/div898/handbook/, 2012.
Horwitz Ratio - Ref: RSC, M Thompson AMC Technical Brief No.17 July 2004
https://www.rsc.org/images/horwitz-function-technica.l-brief-17_tcm18-214859.pdf The "Ideal" Ratio is between 03-1.0. Note that 0.1 is possible for automated sample preparation and handling with modern equipment, but is unlikely for inter-laboratory trials.
Grubbs’ statistic values are at the 0.05 (95%) level for 1-tail. To remove a result, add text (e.g. 123.4?). Removing 2 results from data sets of <7 is not recommended. The method is suspect if outliers are in the same tail.
Ref: RSC - Analytical Methods Committee, AMCTB No. 69
https://pubs.rsc.org/en/content/articlehtml/2015/ay/c5ay90053k and
Frank E. Grubbs and Glenn Beck, "Extension of Sample Sizes and Percentage Points for Significance Tests of Outlying Observations", Technometrics, 14(4), 847-854 (1972).

## Units
Unit | NumericValue
ppm | 1000000
ppb | 1000000000
ppt | 1000000000000
mg/kg | 1000000
µg/kg | 1000000000
ng/kg | 1000000000000
mg/L | 1000000
µg/L | 1000000000
ng/L | 1000000000000
pg/L | 1000000000000000
g/L | 100
percent | 100
drops/floz | 568
If a Unit is added, changed, or deleted - Check that changes to the Unit are in the [MainTable] Units Pop-Up Menu. If there is a mismatch Error Warnings will be displayed…

## Notes
This spreadsheet may be useful to assist Laboratory Analysts and Assessors ensure compliance with ISO/ IEC 17025:2017. The revised standard requires Measurement Uncertainty (MU) of chemical test results to be available. It is designed to help produce necessary information by using sample blanks, spikes, and comparisons with reference materials. Sample "Low Spikes" could be used to determine Recoveries and Method Detection Limits. Spikes establish MUs at different concentrations with different matrices. It may be helpful when developing new methods or for "novel" analytes and matrices.
Created with Apple "Numbers" Version 13, and Microsoft "Excel" 2013, spreadsheet programs - Using only "Cell Functions". No "Code", Macros, or Hyperlinks are used - It should be “Safe” to distribute…
The spreadsheet uses some conventions: In [MainTable] cells where the user is expected to input data have blue text : labels, bold blue indicates that a selection is made from a list. Default look-ups are grey. Checks have [pale backgrounds]; Warnings have [darker backgrounds]; and failure/out-of-range warnings have [red backgrounds]; Replaced MDLs or Significances have [blue backgrounds] Important data is purple.
Getting Started: The sheets are [MainTable], [SampleResults], [Distributions], [Units], and [Notes]. The last three are "look-ups" or information. Data should be added or changed as outlined below…
[MainTable]: Used to set the parameters of the Test:- The Test Name that was used to create the results is a "look-up" from the [SamplesResults] sheet. The Analyte Name is a "look-up" from the [SamplesResults] sheet. Select Report Type from the drop-down - This will produce different a different [Report] for each type. Select Units from the drop-down - This will change the Horwitz Ratio (larger concentrations should give higher precision). Enter 'Spike Value' for each column (Values between 0.00001 and 10,000,000). MDL (From Mean of Previous Low Spikes) - If a value is entered, it will be used instead of the Results calculation. Select Number of Significant figures - Usually 2 - 1 may be useful if near Quantitation limits, 3 for special precision. Horwitz and Outlier fields with coloured backgrounds are not editable, but require consideration - Outliers or errors may well be the cause… Now go to the [SampleResults] Sheet to check/update outliers as below… [SampleResults]: Used to add SampleID and Results. There are a number of columns:-. The First Column is used to Add/Edit SampleID (or Paste from another application). The 1st Top (Header) row must contain the Test Name. Other Header rows must contain the Analyte Name. The other columns are for plain numerical Results (or Paste from another application). Possible result Outliers can be indicated by adding text e.g. Changing 77 to ?77 or 77! Adding text to Outliers will cause it to be ignored by the Horwitz and Grubbs calculations. If pasted results are not plain numbers, cell formatting will be lost. The sheet can have many rows, but normally 3-11 are used (5, or more recommended). [Distributions]: Used to calculate Method Detection Limit (t-test) and Grubbs’ critical value for 0.05 (95%) level for 1-tail. This is not editable. It is recommended that the user consults the NIST and RSC papers, referenced in the text at the bottom of the sheet, for further background information. [Units]: Used to calculate scale of Horwitz values, and in [Report] and [Calculations - Main Table]:- It is normally Locked, but any changes must be reflected in [MU Calculations] "Units" drop-down.
For further information, please consult these References:- Planning and Reporting Method Validation Studies; Supplement to EURACHEM Guide on the Fitness for Purpose of Analytical Methods, First edition 2019 Blanks in Method Validation, Supplement to EURACHEM Guide The Fitness for Purpose of Analytical Methods, First Edition 2019
The password to unlock individual MS Excel sheets and cells is "timstrutt" (no "quotes"). Apple Sheets use Locked "Shapes".
The [MainTable] uses "hidden" rows with a red font - Changing or deleting them will cause breakages...
Warning: The calculations are based upon standard literature methods, but some approximations have been used. See the references for background information.
This spread-sheet is "Public Domain". Anyone is free to copy, modify, publish, use, or distribute it for any purpose, commercial or non-commercial, and by any means. The author is retired; carries out "Pro Bono" work only; and does not accept any responsibility for any loss caused by the use of, or failure to use, this spreadsheet. Please do not remove my name, copyright & version information.
Version: 2026-06-02 - ©Tim Strutt.