## MainTable
Test Name (TopRow, 1stColumn in SampleResults) | ICP-MS | ICP-MS | ICP-MS | ICP-MS | ICP-MS | ICP-MS | ICP-MS | ICP-MS | ICP-MS | ICP-MS | ICP-MS | ICP-MS | ICP-MS | ICP-MS | ICP-MS | ICP-MS | ICP-MS | ICP-MS | ICP-MS | ICP-MS | ICP-MS | ICP-MS
Analyte (TopRow in SampleResults) | Al | As | B | Ba | Be | Cd | Co | Cr | Cu | Fe | Hg | Mn | Mo | Ni | Pb | Sb | Se | Sn | Sr | Ti | V | Zn
Report Type (Spike, Blank, Ref…) | Spike | Spike | Spike | Spike | Spike | Spike | Spike | Spike | Spike | Spike | Spike | Spike | Spike | Spike | Spike | Spike | Spike | Spike | Spike | Spike | Spike | Spike
Units | µg/kg | µg/kg | µg/kg | µg/kg | µg/kg | µg/kg | µg/kg | µg/kg | µg/kg | µg/kg | µg/kg | µg/kg | µg/kg | µg/kg | µg/kg | µg/kg | µg/kg | µg/kg | µg/kg | µg/kg | µg/kg | µg/kg
Spike Concentration | 95.14 | 2.2 | 99 | 90 | 92 | 9 | 95 | 90 | 90 | 90 | 1 | 90 | 93 | 99 | 90 | 93 | 99 | 90 | 93 | 99 | 100 | 50
MDL - Optional (From Mean of Previous Low Spikes) | Al: 11; As: 0.1
Number of Significant figures | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2
Total Number of Results for Analyte | 9 (1 Excluded!) | 9 | 9 | 9 | 9 | 9 | 9 | 9 | 9 | 9 | 9 | 9 | 9 | 9 | 9 | 9 | 9 | 9 | 9 | 9 | 9 | 9
RSD% (CV) | 7.668% | 8.758% | 7.046% | 38.15% | 9.706% | 6.907% | 7.086% | 12.12% | 7.43% | 149.9% | 7.409% | 9.747% | 28.44% | 7.246% | 11.2% | 10.1% | 7.199% | 7.822% | 9.479% | 8.827% | 9.839% | 9.421%
bias as %age (100 - Recovery) | 12.58% | 26.97% | 3.699% | -22.23% | 11.48% | -3.259% | -2.105% | 10.46% | -3.457% | -126% | 16.37% | -8.889% | -18.16% | 16.24% | -9.753% | -9.08% | 16.12% | -11.85% | 12.33% | 8.193% | 15.11% | 15.13%
Horwitz Ratio (Ideally 0.3≤ HorRat ≤1) | 0.34856 | 0.39809 | 0.32026 | 1.7343 | 0.44118 | 0.31396 | 0.32211 | 0.55084 | 0.33773 | 6.813 | 0.33679 | 0.44307 | 1.2927 | 0.32938 | 0.5092 | 0.4592 | 0.32725 | 0.35554 | 0.43085 | 0.40122 | 0.44721 | 0.42821
Outlier? | 0 | 0 | 0 | 220.1 | 0 | 0 | 0 | 59.97 | 0 | 1016 | 0 | 0 | 193 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Result to 2 Significant figures | 83 | 1.6 | 95 | 110 | 81 | 9.3 | 97 | 81 | 93 | 200 | 0.84 | 98 | 110 | 83 | 99 | 100 | 83 | 100 | 82 | 91 | 85 | 42
MDL = (MDL Factor x SD) | 11 | 0.10 | 19 | 120 | 23 | 1.9 | 20 | 28 | 20 | 880 | 0.18 | 28 | 91 | 17 | 32 | 30 | 17 | 23 | 22 | 23 | 24 | 12
PQL (Practical Quantitation limit) | 33 | 0.30 | 58 | 360 | 69 | 5.6 | 60 | 85 | 60 | 2600 | 0.54 | 83 | 270 | 52 | 96 | 89 | 52 | 68 | 67 | 70 | 73 | 35
% Recovery | 87.42% | 73.03% | 96.3% | 122.2% | 88.52% | 103.3% | 102.1% | 89.54% | 103.5% | 226% | 83.63% | 108.9% | 118.2% | 83.76% | 109.8% | 109.1% | 83.88% | 111.9% | 87.67% | 91.81% | 84.89% | 84.87% | 0% | 0% | 0% | 0% | 0% | 0% | 0% | 0% | 0% | 0% | 0% | 0% | 0% | 0%
Result Corrected for Recovery | 95 | 2.2 | 99 | 90 | 92 | 9.0 | 95 | 90 | 90 | 90 | 1.0 | 90 | 93 | 99 | 90 | 93 | 99 | 90 | 93 | 99 | 100 | 50
Ymax (Maximum Value) | 89.84 | 1.82 | 107.83 | 220.1 | 89.84 | 10.4 | 106 | 89.84 | 103 | 1016 | 0.898 | 117 | 193 | 89.84 | 121 | 122 | 89.84 | 117 | 89.84 | 100 | 101.7 | 49.3
Ymin (Minimum Value) | 72.11 | 1.36 | 85.37 | 81.31 | 67.61 | 8.07 | 83 | 59.97 | 84 | 96.1 | 0.721 | 90 | 96 | 72.11 | 90 | 84 | 72.11 | 88 | 68.43 | 77 | 74.7 | 37.6
Mean of Samples (Ymv) | 83.1675 | 1.6067 | 95.3378 | 110.01 | 81.4389 | 9.2933 | 97 | 80.59 | 93.1111 | 203.3667 | 0.83633 | 98 | 109.8889 | 82.9267 | 98.7778 | 101.4444 | 83.0378 | 100.6667 | 81.53 | 90.8889 | 84.8889 | 42.4333
ST DEV σ of Results | 6.3775 | 0.14071 | 6.7172 | 41.9732 | 7.9045 | 0.64189 | 6.8739 | 9.7663 | 6.9182 | 304.8181 | 0.061968 | 9.5525 | 31.2508 | 6.0092 | 11.0655 | 10.2483 | 5.9783 | 7.874 | 7.7279 | 8.0225 | 8.3519 | 3.9975
Uncertainty% at 95% | 0.29473 | 0.56712 | 0.15915 | 0.88319 | 0.30066 | 0.15275 | 0.14785 | 0.32011 | 0.1639 | 3.9157 | 0.35931 | 0.26384 | 0.67484 | 0.35559 | 0.29706 | 0.27167 | 0.35316 | 0.28401 | 0.3111 | 0.24086 | 0.36063 | 0.35652
MDL Factor (n-1 Degrees Freedom) | 2.998 | 2.896 | 2.896 | 2.896 | 2.896 | 2.896 | 2.896 | 2.896 | 2.896 | 2.896 | 2.896 | 2.896 | 2.896 | 2.896 | 2.896 | 2.896 | 2.896 | 2.896 | 2.896 | 2.896 | 2.896 | 2.896
Units - 1 Part in This Value | 1000000000 | 1000000000 | 1000000000 | 1000000000 | 1000000000 | 1000000000 | 1000000000 | 1000000000 | 1000000000 | 1000000000 | 1000000000 | 1000000000 | 1000000000 | 1000000000 | 1000000000 | 1000000000 | 1000000000 | 1000000000 | 1000000000 | 1000000000 | 1000000000 | 1000000000
Predicted Horwitz "Expected Value" as RSD% | 22 | 22 | 22 | 22 | 22 | 22 | 22 | 22 | 22 | 22 | 22 | 22 | 22 | 22 | 22 | 22 | 22 | 22 | 22 | 22 | 22 | 22
Number of Results included in Calculations | 8 | 9 | 9 | 9 | 9 | 9 | 9 | 9 | 9 | 9 | 9 | 9 | 9 | 9 | 9 | 9 | 9 | 9 | 9 | 9 | 9 | 9
Ymax - Ymv | 6.6725 | 0.21333 | 12.4922 | 110.09 | 8.4011 | 1.1067 | 9 | 9.25 | 9.8889 | 812.6333 | 0.061667 | 19 | 83.1111 | 6.9133 | 22.2222 | 20.5556 | 6.8022 | 16.3333 | 8.31 | 9.1111 | 16.8111 | 6.8667
Ymv - Ymin | 11.0575 | 0.24667 | 9.9678 | 28.7 | 13.8289 | 1.2233 | 14 | 20.62 | 9.1111 | 107.2667 | 0.11533 | 8 | 13.8889 | 10.8167 | 8.7778 | 17.4444 | 10.9278 | 12.6667 | 13.1 | 13.8889 | 10.1889 | 4.8333
Sig Value: 1-sided | 0.0125 | 0.011111 | 0.011111 | 0.011111 | 0.011111 | 0.011111 | 0.011111 | 0.011111 | 0.011111 | 0.011111 | 0.011111 | 0.011111 | 0.011111 | 0.011111 | 0.011111 | 0.011111 | 0.011111 | 0.011111 | 0.011111 | 0.011111 | 0.011111 | 0.011111
G (min) | 1.7338 | 1.753 | 1.4839 | 0.68377 | 1.7495 | 1.9058 | 2.0367 | 2.1114 | 1.317 | 0.3519 | 1.8612 | 0.83748 | 0.44443 | 1.8 | 0.79326 | 1.7022 | 1.8279 | 1.6087 | 1.6952 | 1.7312 | 1.22 | 1.2091
G (max) | 1.0463 | 1.5161 | 1.8597 | 2.6229 | 1.0628 | 1.7241 | 1.3093 | 0.94714 | 1.4294 | 2.666 | 0.99514 | 1.989 | 2.6595 | 1.1505 | 2.0083 | 2.0058 | 1.1378 | 2.0743 | 1.0753 | 1.1357 | 2.0129 | 1.7177
G Critical | 2.032 | 2.11 | 2.11 | 2.11 | 2.11 | 2.11 | 2.11 | 2.11 | 2.11 | 2.11 | 2.11 | 2.11 | 2.11 | 2.11 | 2.11 | 2.11 | 2.11 | 2.11 | 2.11 | 2.11 | 2.11 | 2.11
Check if This Result is an Outlier | 0 | 0 | 0 | 220.1 | 0 | 0 | 0 | 59.97 | 0 | 1016 | 0 | 0 | 193 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Number of Excluded Results | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
uncertainty at 95Conf - u | 6.7643 | 0.14832 | 7.0805 | 44.2436 | 8.3321 | 0.67661 | 7.2457 | 10.2945 | 7.2924 | 321.3065 | 0.06532 | 10.0692 | 32.9412 | 6.3342 | 11.664 | 10.8027 | 6.3016 | 8.2999 | 8.1459 | 8.4565 | 8.8036 | 4.2137
MDL = (MDL Factor x SD) | 11 | 0.1 | 19.453 | 121.5544 | 22.8914 | 1.8589 | 19.9067 | 28.2831 | 20.035 | 882.7531 | 0.17946 | 27.664 | 90.5023 | 17.4026 | 32.0456 | 29.6791 | 17.313 | 22.8031 | 22.3801 | 23.2333 | 24.187 | 11.5768
PLQ | 33 | 0.3 | 58.3589 | 364.6632 | 68.6742 | 5.5768 | 59.7201 | 84.8492 | 60.1051 | 2648.2593 | 0.53838 | 82.992 | 271.5068 | 52.2077 | 96.1367 | 89.0373 | 51.9391 | 68.4094 | 67.1402 | 69.6998 | 72.561 | 34.7303
Bias | 13.5287 | 0.29665 | 14.1611 | 88.4873 | 16.6641 | 1.3532 | 14.4914 | 20.5891 | 14.5848 | 642.6129 | 0.13064 | 20.1384 | 65.8824 | 12.6685 | 23.328 | 21.6053 | 12.6033 | 16.5999 | 16.2919 | 16.913 | 17.6073 | 8.4275
Raw bias (Mean - Spike) | -11.9725 | -0.59333 | -3.6622 | 20.01 | -10.5611 | 0.29333 | 2 | -9.41 | 3.1111 | 113.3667 | -0.16367 | 8 | 16.8889 | -16.0733 | 8.7778 | 8.4444 | -15.9622 | 10.6667 | -11.47 | -8.1111 | -15.1111 | -7.5667
MDL Corrected | 12.5835 | 0.13693 | 20.2002 | 99.4446 | 25.86 | 1.8002 | 19.4963 | 31.5855 | 19.3656 | 390.6627 | 0.21458 | 25.4057 | 76.5929 | 20.7756 | 29.1979 | 27.2085 | 20.6411 | 20.3869 | 25.5286 | 25.3067 | 28.4925 | 13.6411
MDL rounded | 11 | 0.10 | 19 | 120 | 23 | 1.9 | 20 | 28 | 20 | 880 | 0.18 | 28 | 91 | 17 | 32 | 30 | 17 | 23 | 22 | 23 | 24 | 12
PQL rounded | 33 | 0.30 | 58 | 360 | 69 | 5.6 | 60 | 85 | 60 | 2600 | 0.54 | 83 | 270 | 52 | 96 | 89 | 52 | 68 | 67 | 70 | 73 | 35
Spike Level to correct Sig. Figs | 95 | 2.2 | 99 | 90 | 92 | 9.0 | 95 | 90 | 90 | 90 | 1.0 | 90 | 93 | 99 | 90 | 93 | 99 | 90 | 93 | 99 | 100 | 50
MDL Corrected rounded | 13 | 0.14 | 20 | 100 | 26 | 1.8 | 19 | 32 | 19 | 390 | 0.21 | 25 | 77 | 21 | 29 | 27 | 21 | 20 | 26 | 25 | 28 | 14
Corrected for Recovery with Uncertainy | 95 ± 13 µg/kg | 2.2 ± 0.14 µg/kg | 99 ± 20 µg/kg | 90 ± 100 µg/kg | 92 ± 26 µg/kg | 9.0 ± 1.8 µg/kg | 95 ± 19 µg/kg | 90 ± 32 µg/kg | 90 ± 19 µg/kg | 90 ± 390 µg/kg | 1.0 ± 0.21 µg/kg | 90 ± 25 µg/kg | 93 ± 77 µg/kg | 99 ± 21 µg/kg | 90 ± 29 µg/kg | 93 ± 27 µg/kg | 99 ± 21 µg/kg | 90 ± 20 µg/kg | 93 ± 26 µg/kg | 99 ± 25 µg/kg | 100 ± 28 µg/kg | 50 ± 14 µg/kg

## SampleResults
ICP-MS | Al | As | B | Ba | Be | Cd | Co | Cr | Cu | Fe | Hg | Mn | Mo | Ni | Pb | Sb | Se | Sn | Sr | Ti | V | Zn
S1 | 86.38 | 1.59 | 90.21 | 94.21 | 86.38 | 9.03 | 96 | 86.38 | 90 | 96.1 | 0.851 | 90 | 96 | 86.38 | 90 | 84 | 86.38 | 96 | 86.38 | 90 | 82.1 | 41.2
S2 | 87.79 | 1.66 | 91.3 | 88.46 | 87.79 | 9.11 | 97 | 87.79 | 84 | 97.2 | 0.882 | 91 | 97 | 87.79 | 91 | 97 | 87.79 | 97 | 87.79 | 91 | 79.1 | 39.7
S3 | 77.94 | 1.42 | 91.88 | 98.27 | 77.94 | 9.24 | 98 | 77.94 | 101 | 98.2 | 0.779 | 92 | 98 | 77.94 | 92 | 98 | 77.94 | 88 | 77.94 | 79 | 85.6 | 43.6
S4 | 72.11 | 1.59 | 107.83 | 99.41 | 72.11 | 9.35 | 83 | 72.11 | 93 | 99.3 | 0.721 | 93 | 99 | 72.11 | 93 | 99 | 72.11 | 99 | 72.11 | 93 | 88.7 | 44.5
S5 Duplicate S3 | 78.56 | 1.66 | 94.69 | 100.24 | 78.56 | 9.06 | 100 | 78.56 | 85 | 100.2 | 0.786 | 94 | 100 | 78.56 | 94 | 100 | 78.56 | 100 | 78.56 | 94 | 92.3 | 47.1
S6 | 89.84 | 1.36 | 85.37 | 100.91 | 89.84 | 9.47 | 101 | 89.84 | 95 | 101.3 | 0.898 | 95 | 101 | 89.84 | 95 | 101 | 89.84 | 101 | 89.84 | 95 | 101.7 | 49.3
S7 | 89.3 | 1.69 | 99.21 | 81.31 | 89.3 | 9.91 | 102 | 89.3 | 99 | 102.4 | 0.893 | 99 | 102 | 89.3 | 99 | 102 | 89.3 | 102 | 89.3 | 99 | 77.6 | 38.2
S8 | 83.42 | 1.82 | 96.33 | 107.18 | 83.42 | 8.07 | 90 | 83.42 | 103 | 1016 | 0.833 | 117 | 103 | 83.42 | 121 | 110 | 83.42 | 106 | 83.42 | 100 | 82.2 | 40.7
S9 | !58.97 | 1.67 | 101.22 | 220.1 | 67.61 | 10.4 | 106 | 59.97 | 88 | 119.6 | 0.884 | 111 | 193 | 81 | 114 | 122 | 82 | 117 | 68.43 | 77 | 74.7 | 37.6

## Distributions
Student’s t-test |  | Grubbs’ test
Limit | 99% | 𝜶 0.05 1-sided
1 | 31.82
2 | 6.965
3 | 4.541 | 1.153
4 | 3.747 | 1.463
5 | 3.365 | 1.672
6 | 3.143 | 1.822
7 | 2.998 | 1.938
8 | 2.896 | 2.032
9 | 2.821 | 2.11
10 | 2.764 | 2.176
11 | 2.718 | 2.234
12 | 2.681 | 2.285
13 | 2.65 | 2.331
14 | 2.624 | 2.371
15 | 2.602 | 2.409
16 | 2.583 | 2.443
17 | 2.567 | 2.475
18 | 2.552 | 2.504
19 | 2.539 | 2.532
20 | 2.528 | 2.557
21 | 2.518 | 2.58
22 | 2.508 | 2.603
23 | 2.5 | 2.624
24 | 2.492 | 2.644
25 | 2.485 | 2.663
26 | 2.479 | 2.681
27 | 2.473 | 2.698
28 | 2.467 | 2.714
29 | 2.462 | 2.73
30 | 2.457 | 2.745
40 | 2.423 | 2.866
50 | 2.403 | 2.956
60 | 2.39 | 3.025
80 | 2.374 | 3.13
100 | 2.364 | 3.207
150 | 2.352 | 3.35
200 | 2.345 | 3.5
1000 | 2.33 | 4
Student’s t-distributions have ν degrees of freedom for regions.
See www.itl.nist.gov/div898/handbook/, 2012.
Horwitz Ratio - Ref: RSC, M Thompson AMC Technical Brief No.17 July 2004
https://www.rsc.org/images/horwitz-function-technica.l-brief-17_tcm18-214859.pdf The "Ideal" Ratio is between 03-1.0. Note that 0.1 is possible for automated sample preparation and handling with modern equipment, but is unlikely for inter-laboratory trials.
Grubbs’ statistic values are at the 0.05 (95%) level for 1-tail. To remove a result, add text (e.g. 123.4?). Removing 2 results from data sets of <7 is not recommended. The method is suspect if outliers are in the same tail.
Ref: RSC - Analytical Methods Committee, AMCTB No. 69
https://pubs.rsc.org/en/content/articlehtml/2015/ay/c5ay90053k and
Frank E. Grubbs and Glenn Beck, "Extension of Sample Sizes and Percentage Points for Significance Tests of Outlying Observations", Technometrics, 14(4), 847-854 (1972).

## Units
Unit | NumericValue
ppm | 1000000
ppb | 1000000000
ppt | 1000000000000
mg/kg | 1000000
µg/kg | 1000000000
ng/kg | 1000000000000
mg/L | 1000000
µg/L | 1000000000
ng/L | 1000000000000
pg/L | 1000000000000000
g/L | 100
percent | 100
drops/floz | 568
If a Unit is added, changed, or deleted - Check that changes to the Unit are in the [MainTable] Units Pop-Up Menu. If there is a mismatch Error Warnings will be displayed…

## Notes
This spreadsheet may be useful to assist Laboratory Analysts and Assessors ensure compliance with ISO/ IEC 17025:2017. The revised standard requires Measurement Uncertainty (MU) of chemical test results to be available. It is designed to help produce necessary information by using sample blanks, spikes, and comparisons with reference materials. Sample "Low Spikes" could be used to determine Recoveries and Method Detection Limits. Spikes establish MUs at different concentrations with different matrices. It may be helpful when developing new methods or for "novel" analytes and matrices.
Created with Apple "Numbers" Version 12.0, and Microsoft "Excel" 2013, spreadsheet programs - Using only "Cell Functions". No "Code", Macros, or Hyperlinks are used - It should be “Safe” to distribute…
The spreadsheet uses some conventions: In [MainTable] cells where the user is expected to input data have blue text on a white background, bold blue indicates that a selection is made from a list. Default look-ups are grey. Warnings are highlighted with an [orange background]; and critical/failure/out-of-range warnings are highlighted with [red backgrounds]. Important data is purple.
Getting Started: The sheets are [MainTable], [SampleResults], [Distributions], [Units], and [Notes]. The last three are "look-ups" or information. Data should be added or changed as outlined below…
[MainTable]: Used to set the parameters of the Test:- The Test Name that was used to create the results is a "look-up" from the [SamplesResults] sheet. The Analyte Name is a "look-up" from the [SamplesResults] sheet. Select Report Type from the drop-down - This will produce different a different [Report] for each type. Select Units from the drop-down - This will change the Horwitz Ratio (larger concentrations should give higher precision). Enter 'Spike Value' for each column (Values between 0.00001 and 10,000,000). MDL (From Mean of Previous Low Spikes) - If a value is entered, it will be used instead of the Results calculation. Select Number of Significant figures - Usually 2 - 1 may be useful if near Quantitation limits, 3 for special precision. Horwitz and Outlier fields with coloured backgrounds are not editable, but require consideration - Outliers or errors may well be the cause… Now go to the [SampleResults] Sheet to check/update outliers as below… [SampleResults]: Used to add SampleID and Results. There are a number of columns:-. The First Column is used to Add/Edit SampleID (or Paste from another application). The 1st Top (Header) row must contain the Test Name. Other Header rows must contain the Analyte Name. The other columns are for plain numerical Results (or Paste from another application). Possible result Outliers can be indicated by adding text e.g. Changing 77 to ?77 or 77! Adding text to Outliers will cause it to be ignored by the Horwitz and Grubbs calculations. If pasted results are not plain numbers, cell formatting will be lost. The sheet can have many rows, but normally 3-11 are used (5, or more recommended). [Distributions]: Used to calculate Method Detection Limit (t-test) and Grubbs’ critical value for 0.05 (95%) level for 1-tail. This is not editable. It is recommended that the user consults the NIST and RSC papers, referenced in the text at the bottom of the sheet, for further background information. [Units]: Used to calculate scale of Horwitz values, and in [Report] and [Calculations - Main Table]:- It is normally Locked, but any changes must be reflected in [MU Calculations] "Units" drop-down.
For further information, please consult these References:- Planning and Reporting Method Validation Studies; Supplement to EURACHEM Guide on the Fitness for Purpose of Analytical Methods, First edition 2019 Blanks in Method Validation, Supplement to EURACHEM Guide The Fitness for Purpose of Analytical Methods, First Edition 2019
The password to unlock individual MS Excel sheets and cells is "timstrutt" (no "quotes"). Apple Sheets use Locked "Shapes".
The [MainTable] uses "hidden" rows with a red font - Changing or deleting them will cause breakages...
Warning: The calculations are based upon standard literature methods, but some approximations have been used. See the references for background information.
This spread-sheet is "Public Domain". Anyone is free to copy, modify, publish, use, or distribute it for any purpose, commercial or non-commercial, and by any means. The author is retired; carries out "Pro Bono" work only; and does not accept any responsibility for any loss caused by the use of, or failure to use, this spreadsheet. Please do not remove my name, copyright & version information.
Version: 2022-10-20 - ©Tim Strutt